## Data
Centre_ID | Group | CentreType | Date | Type | Energy | CoP | BeamQualRatio | BeamQualError | OPRatio | OPError | IntercompRatio | IntercompError | Comments
14 | 2 | NHS | 1994-06-07 | MV | 8 |  | 0.98579 | 0.2 | 1.006 | 0.2 | 0.99925 | 0.4
40 | 4 | NHS | 1994-07-27 | MV | 6 |  | 1.0025 | 0.2 | 1.0062 | 0.2
1 | 1 | NHS | 2007-02-10 | MeV | 4 | 2003 | 0.01 | 0.2 | 0.998 | 0.2 | 1.001 | 0.4
1 | 1 | NHS | 2007-02-10 | MeV | 6 | 2003 | 0.02 | 0.2 | 1.003 | 0.2 | 1.002 | 0.4
1 | 1 | NHS | 2007-02-10 | MeV | 9 | 2003 | 0.03 | 0.2 | 1.003 | 0.2 | 1.001 | 0.4
1 | 1 | NHS | 2007-02-10 | MeV | 12 | 2003 | 0.05 | 0.2 | 0.999 | 0.2 | 1.001 | 0.4
1 | 1 | NHS | 2007-02-10 | MeV | 16 | 2003 | -0.01 | 0.2 | 1.002 | 0.2 | 1.002 | 0.4
40 | 4 | NHS | 1998-06-11 | kV | 30 |  |  |  | 1.0119 | 0.6
6 | 6 | NHS | 2005-08-09 | kV | 85 |  | 1.07 | 3 | 0.999 | 0.6
6 | 6 | NHS | 2005-08-09 | kV | 250 |  | 1.02 | 5 | 1.008 | 0.8 | 1.001 | 0.4
26 | 3 | NHS | 1994-10-18 | MV | 6 |  | 0.9937 | 0.2 | 1.0013 | 0.2 | 1.009 | 0.4
7 | 6 | NHS | 1994-11-16 | MV | 8 |  | 1.0006 | 0.2 | 0.996 | 0.2 | 0.99434 | 0.4
22 | 7 | NHS | 1995-06-15 | MV | 6 |  | 1.0103 | 0.2 | 1.0031 | 0.2 | 1.0032 | 0.4
3 | 1 | NHS | 1995-06-21 | MV | 6 |  | 0.99603 | 0.2 | 0.98803 | 0.2 | 0.98584 | 0.4
39 | 4 | NHS | 1996-04-24 | MV | 15 |  | 0.9987 | 0.2 | 1.0083 | 0.2 | 0.99555 | 0.4
44 | 5 | NHS | 1996-05-15 | MV | 16 |  | 1.0066 | 0.2 | 1.0111 | 0.2 | 1.0109 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
4 | 1 | NHS | 1996-06-19 | MV | 6 |  | 0.9941 | 0.2 | 1.0019 | 0.2 | 1.0069 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
10 | 6 | NHS | 1996-06-21 | MV | 8 |  | 1.0042 | 0.2 | 1.0199 | 0.2 | 1.0043 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
20 | 7 | NHS | 1996-09-03 | MV | 6 |  | 0.997 | 0.2 | 1.007 | 0.2 | 1.0005 | 0.4
14 | 2 | NHS | 1996-10-10 | MV | 8 |  | 0.985 | 0.2 | 1.001 | 0.2 | 0.9995 | 0.4
13 | 2 | NHS | 1996-10-31 | MV | 6 |  | 1.0013 | 0.2 | 1.0194 | 0.2 | 1.0018 | 0.4
13 | 2 | NHS | 1996-10-31 | MV | 25 |  | 0.98707 | 0.2 | 1.0124 | 0.2 | 1.0097 | 0.4
38 | 4 | NHS | 1997-02-20 | MV | 5 |  | 0.9939 | 0.2 | 1.009 | 0.2 | 1.003 | 0.4
33 | 3 | NHS | 1997-06-17 | MV | 6 |  | 1 | 0.2 | 0.991 | 0.2 | 0.984 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
46 | 5 | NHS | 2000-03-21 | MV | 6 |  | 0.993 | 0.15 | 0.996 | 0.2 | 0.958 | 0.4 | Looks like a preliminary report
5 | 6 | NHS | 2000-03-24 | MV | 6 |  | 1.004 | 0.2 | 0.991 | 0.2 | 1.004 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
18 | 7 | NHS | 2000-05-16 | MV | 6 |  | 1 | 0.2 | 1.005 | 0.2 | 0.999 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
3 | 1 | NHS | 2008-10-14 | MeV | 6 | 2003 | -0.06 | 0.2 | 0.996 | 0.2 | 1.003 | 0.4
3 | 1 | NHS | 2008-10-14 | MeV | 9 | 2003 | -0.08 | 0.2 | 1.015 | 0.2 | 1.003 | 0.4
3 | 1 | NHS | 2008-10-14 | MeV | 12 | 2003 | -0.07 | 0.2 | 1.024 | 0.2 | 1.013 | 0.4
3 | 1 | NHS | 2008-10-14 | MeV | 16 | 2003 | -0.16 | 0.2 | 1.021 | 0.2 | 0.994 | 0.4
3 | 1 | NHS | 2008-10-14 | MeV | 20 | 2003 | -0.27 | 0.2 | 1.027 | 0.2 | 0.996 | 0.4
3 | 1 | NHS | 2011-05-23 | MeV | 6 | 2003 | -0.19 | 0.2 | 0.995 | 0.2 | 0.996 | 0.4
3 | 1 | NHS | 2011-05-23 | MeV | 9 | 2003 | -0.09 | 0.2 | 0.995 | 0.2 | 0.997 | 0.4
3 | 1 | NHS | 2011-05-23 | MeV | 12 | 2003 | -0.09 | 0.2 | 0.996 | 0.2 | 0.997 | 0.4
3 | 1 | NHS | 2011-05-23 | MeV | 16 | 2003 | -0.2 | 0.2 | 0.985 | 0.2 | 0.987 | 0.4
3 | 1 | NHS | 2011-05-23 | MeV | 20 | 2003 | -0.28 | 0.2 | 0.987 | 0.2 | 0.99 | 0.4
3 | 1 | NHS | 2013-08-28 | MeV | 6 | 2003 | -0.06 | 0.2 | 1.004 | 0.2 | 0.999 | 0.4
3 | 1 | NHS | 2013-08-28 | MeV | 9 | 2003 | -0.06 | 0.2 | 0.998 | 0.2 | 0.997 | 0.4
3 | 1 | NHS | 2013-08-28 | MeV | 12 | 2003 | -0.13 | 0.2 | 0.998 | 0.2 | 0.997 | 0.4
3 | 1 | NHS | 2013-08-28 | MeV | 16 | 2003 | -0.04 | 0.2 | 0.993 | 0.2 | 1.001 | 0.4
3 | 1 | NHS | 2013-08-28 | MeV | 20 | 2003 | -0.13 | 0.2 | 1.007 | 0.2 | 1.011 | 0.4
21 | 7 | NHS | 2005-08-25 | kV | 60 |  |  |  | 1.002 | 0.6
21 | 7 | NHS | 2005-08-25 | kV | 80 |  |  |  | 0.976 | 0.6
18 | 7 | NHS | 2000-05-16 | MV | 10 |  | 1.005 | 0.2 | 1 | 0.2 | 0.991 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
3 | 1 | NHS | 2000-06-13 | MV | 6 |  | 1 | 0.2 | 1.004 | 0.2 | 1.002 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
17 | 2 | NHS | 2000-06-14 | MV | 6 |  | 0.998 | 0.2 | 0.999 | 0.2 | 0.997 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
4 | 1 | NHS | 2008-09-16 | MeV | 6 | 2003 | 0.03 | 0.2 | 1.011 | 0.2
4 | 1 | NHS | 2008-09-16 | MeV | 9 | 2003 | -0.03 | 0.2 | 1.004 | 0.2
4 | 1 | NHS | 2008-09-16 | MeV | 12 | 2003 | -0.05 | 0.2 | 1.012 | 0.2
4 | 1 | NHS | 2008-09-16 | MeV | 18 | 2003 | -0.05 | 0.2 | 1.003 | 0.2
17 | 2 | NHS | 2000-06-14 | MV | 6 |  | 0.998 | 0.2 | 0.999 | 0.2 | 0.997 | 0.4
37 | 4 | NHS | 2000-12-14 | MV | 6 |  | 0.9963 | 0.2 | 1.007 | 0.2 | 1.0053 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
29 | 3 | NHS | 2001-03-22 | MV | 10 |  | 1.0045 | 0.2 | 0.9989 | 0.2 | 1.0038 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
6 | 6 | NHS | 2005-08-09 | MeV | 6 | 2003 |  |  | 1 | 0.2 |  |  | Used host supplied R50D
6 | 6 | NHS | 2005-08-09 | MeV | 9 | 2003 |  |  | 1.001 | 0.2 |  |  | Used host supplied R50D
6 | 6 | NHS | 2005-08-09 | MeV | 16 | 2003 |  |  | 1.008 | 0.2 |  |  | Used host supplied R50D
2 | 1 | NHS | 2007-02-08 | kV | 90 |  | 1.07 | 2 | 1.006 | 0.6 | 0.994 | 0.4
2 | 1 | NHS | 2007-02-08 | kV | 135 |  | 1.02 | 2 | 1.009 | 0.6 | 1 | 0.4
44 | 5 | NHS | 2001-06-04 | MV | 18 |  | 1 | 0.2 | 1.003 | 0.2 | 1.007 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
34 | 8 | NHS | 2001-09-29 | MV | 6 |  | 0.998 | 0.2 | 1.01 | 0.2 | 1.002 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
9 | 6 | NHS | 2002-01-21 | MV | 6 |  | 0.997 | 0.2 | 1.007 | 0.2 | 1.005 | 0.4
9 | 6 | NHS | 2002-01-21 | MV | 10 |  | 0.996 | 0.2 | 1.007 | 0.2 | 1.007 | 0.4
45 | 5 | NHS | 2002-12-18 | MV | 6 |  | 1.003 | 0.15 | 0.987 | 0.1 | 0.999 | 0.4 | Two field instruments checked. One had ratio of 1.000, other 0.999
9 | 6 | NHS | 2001-09-29 | MeV | 6 | 1996 |  |  | 1.0165 | 0.2
9 | 6 | NHS | 2001-09-29 | MeV | 12 | 1996 |  |  | 1.0097 | 0.2
45 | 5 | NHS | 2002-12-18 | MV | 10 |  | 0.994 | 0.15 | 0.992 | 0.1 | 1.001 | 0.4 | Two field instruments checked. Both had ratio of 1.001
19 | 7 | NHS | 2003-01-13 | MV | 4 |  | 1.002 | 0.2 | 0.999 | 0.2 | 0.999 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
19 | 7 | NHS | 2003-01-13 | MV | 6 |  | 1.007 | 0.2 | 1.003 | 0.2 | 1.002 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
12 | 2 | NHS | 2000-11-24 | MeV | 6 | 1996 |  |  | 1.004 | 0.5
12 | 2 | NHS | 2000-11-24 | MeV | 10 | 1996 |  |  | 1.001 | 0.5
19 | 7 | NHS | 2003-01-13 | MV | 10 |  | 1.003 | 0.2 | 1.003 | 0.2 | 1.002 | 0.4 | Host data not given in spreadsheet. Therefore could not calculate ratios. Data abtained from summary produced by RAST.
6 | 6 | NHS | 2005-08-09 | MV | 6 |  | 1 | 0.15 | 1.005 | 0.1 | 1.006 | 0.4 | Using half voltage method to get correction factor of 1.010. Usign formula for PRF=72Hz gave factor of 1.023. This led to output ratio of 1.017 and intercomp of 1.018
13 | 2 | NHS | 2007-03-23 | MeV | 4 | 2003 | -0.02 |  | 1.012 | 0.2 | 1.008 | 0.4
13 | 2 | NHS | 2007-03-23 | MeV | 6 | 2003 | -0.08 |  | 1.009 | 0.2 | 1.006 | 0.4
13 | 2 | NHS | 2007-03-23 | MeV | 10 | 2003 | 0.12 |  | 1.01 | 0.2 | 1.007 | 0.4
13 | 2 | NHS | 2007-03-23 | MeV | 15 | 2003 | 0.12 |  | 1.01 | 0.2 | 1.007 | 0.4
13 | 2 | NHS | 2007-03-23 | MeV | 20 | 2003 | 0.08 |  | 1.001 | 0.2 | 1.003 | 0.4
6 | 6 | NHS | 2005-08-09 | MV | 10 |  | 1 | 0.15 | 1.003 | 0.1 | 1.003 | 0.4 | Using half voltage method to get correction factor of 1.008. Usign formula for PRF=72Hz gave factor of 1.023. This led to output ratio of 1.018 and intercomp of 1.018
44 | 5 | NHS | 2005-10-11 | MV | 6 |  | 0.999 | 0.15 | 0.995 | 0.1 | 1 | 0.4
2 | 1 | NHS | 2007-02-08 | kV | 300 |  | 1 | 2 | 0.995 | 0.6 | 0.997 | 0.4
31 | 3 | NHS | 2007-02-12 | kV | 60 |  | 0.996 | 2 | 0.987 | 0.6 | 1.011 | 0.4
31 | 3 | NHS | 2007-02-12 | kV | 100 |  | 0.993 | 2 | 1.001 | 0.6 | 1.002 | 0.4
31 | 3 | NHS | 2007-02-12 | kV | 160 |  | 1.005 | 2 | 0.997 | 0.6 | 1 | 0.4
15 | 2 | NHS | 2010-04-30 | MeV | 6 | 2003 | -0.02 |  | 0.999 | 0.2 | 0.999 | 0.4
15 | 2 | NHS | 2010-04-30 | MeV | 9 | 2003 | -0.03 |  | 0.997 | 0.2 | 0.997 | 0.4
15 | 2 | NHS | 2010-04-30 | MeV | 12 | 2003 | -0.03 |  | 1 | 0.2 | 1 | 0.4
15 | 2 | NHS | 2010-04-30 | MeV | 15 | 2003 | -0.01 |  | 1.002 | 0.2 | 1.002 | 0.4
44 | 5 | NHS | 2005-10-11 | MV | 18 |  | 0.998 | 0.15 | 1.002 | 0.2 | 1.001 | 0.4
23 | 7 | NHS | 2005-11-18 | MV | 6 |  | 0.998 | 0.2 | 0.992 | 0.2 | 0.997 | 0.4
23 | 7 | NHS | 2005-11-18 | MV | 10 |  | 0.997 | 0.2 | 0.992 | 0.2 | 1 | 0.4
31 | 3 | NHS | 2006-10-25 | MV | 6 |  | 0.995 | 0.15 | 1.017 | 0.15 | 1.005 | 0.4
31 | 3 | NHS | 2006-10-25 | MV | 10 |  | 0.999 | 0.15 | 1.01 | 0.15 | 1.003 | 0.4
24 | 3 | NHS | 2007-02-02 | MV | 6 |  | 0.993 | 0.15 | 0.999 | 0.1 | 0.999 | 0.4 | Measured Values
18 | 7 | NHS | 2000-09-28 | MeV | 9 | 1996 |  |  | 1.01 | 0.5
18 | 7 | NHS | 2000-09-28 | MeV | 12 | 1996 |  |  | 1.004 | 0.5
24 | 3 | NHS | 2007-02-02 | MV | 15 |  | 0.997 | 0.15 | 1 | 0.1 | 1.002 | 0.4 | Measured Values
1 | 1 | NHS | 2007-02-09 | MV | 6 |  | 0.993 | 0.15 | 1.005 | 0.1 | 0.992 | 0.4 | Measured result (1.003 from tabulated data)
1 | 1 | NHS | 2007-02-09 | MV | 15 |  | 0.995 | 0.15 | 1.008 | 0.1 | 0.991 | 0.4
31 | 3 | NHS | 2007-02-12 | kV | 220 |  | 0.912 | 2 | 1.004 | 0.6 | 0.992 | 0.4
14 | 2 | NHS | 2007-04-23 | kV | 80 |  | 0.963 | 2 | 1.001 | 0.6 | 0.996 | 0.4
14 | 2 | NHS | 2007-04-23 | kV | 120 |  | 0.979 | 2 | 1.008 | 0.6 | 1.006 | 0.4
11 | 2 | NHS | 2007-03-06 | MV | 10 |  | 0.994 | 0.15 | 1 | 0.2 | 1 | 0.4
14 | 2 | NHS | 2007-04-23 | kV | 300 |  | 0.994 | 2 | 0.995 | 0.6 | 1.003 | 0.4 | In Air
14 | 2 | NHS | 2007-04-23 | kV | 300 |  |  |  | 1.006 | 0.6 |  |  | In Water
11 | 2 | NHS | 2007-03-06 | MV | 15 |  | 0.994 | 0.15 | 1.009 | 0.2 | 1.01 | 0.4
24 | 3 | NHS | 2008-08-28 | MV | 6 |  | 0.993 | 0.2 | 0.999 | 0.1 | 1.001 | 0.4
24 | 3 | NHS | 2008-08-28 | MV | 10 |  | 0.996 | 0.2 | 1.001 | 0.1 | 0.997 | 0.4
23 | 7 | NHS | 2005-11-18 | MeV | 6 | 2003 |  |  | 1.006 | 0.2 |  |  | A number of ratios exist in the document with values calculated from 196 and 2003 COPs
23 | 7 | NHS | 2005-11-18 | MeV | 9 | 2003 |  |  | 1.005 | 0.2 |  |  | A number of ratios exist in the document with values calculated from 196 and 2003 COPs
23 | 7 | NHS | 2005-11-18 | MeV | 12 | 2003 |  |  | 1.012 | 0.2 |  |  | A number of ratios exist in the document with values calculated from 196 and 2003 COPs
23 | 7 | NHS | 2005-11-18 | MeV | 15 | 2003 |  |  | 0.998 | 0.2 |  |  | A number of ratios exist in the document with values calculated from 196 and 2003 COPs
23 | 7 | NHS | 2005-11-18 | MeV | 18 | 2003 |  |  | 0.998 | 0.2 |  |  | A number of ratios exist in the document with values calculated from 196 and 2003 COPs
4 | 1 | NHS | 2008-09-16 | MV | 6 |  | 0.991 | 0.2 | 0.993 | 0.2 | 1.002 | 0.4 | Host data is tabulated
4 | 1 | NHS | 2008-09-16 | MV | 15 |  | 0.998 | 0.2 | 1.006 | 0.2 | 1.001 | 0.4 | Host data is tabulated
3 | 1 | NHS | 2008-10-14 | MV | 6 |  | 1.005 | 0.2 | 0.996 | 0.2 | 0.993 | 0.4
3 | 1 | NHS | 2008-10-14 | MV | 10 |  | 1.003 | 0.2 | 0.992 | 0.2 | 0.997 | 0.4
28 | 3 | NHS | 2009-05-19 | MV | 6 |  | 1.012 | 0.15 | 0.997 | 0.2 | 0.991 | 0.4
28 | 3 | NHS | 2009-05-19 | MV | 10 |  | 1.007 | 0.15 | 0.994 | 0.2 | 0.988 | 0.4
25 | 3 | NHS | 2008-08-27 | MeV | 6 | 2003 | 0 |  | 1.001 | 0.2
25 | 3 | NHS | 2008-08-27 | MeV | 9 | 2003 | -0.05 |  | 1.01 | 0.2
25 | 3 | NHS | 2008-08-27 | MeV | 12 | 2003 | 0 |  | 1.006 | 0.2
25 | 3 | NHS | 2008-08-27 | MeV | 16 | 2003 | -0.1 |  | 1.01 | 0.2
25 | 3 | NHS | 2008-08-27 | MeV | 20 | 2003 | 0 |  | 1.008 | 0.2
16 | 2 | NHS | 2009-06-24 | MV | 6 |  | 0.998 | 0.2 | 1.001 | 0.2 | 0.995 | 0.4
16 | 2 | NHS | 2009-06-24 | MV | 10 |  | 1 | 0.2 | 1.001 | 0.2 | 0.993 | 0.4
27 | 3 | NHS | 2009-09-10 | MV | 6 |  | 1.012 | 0.2 | 0.997 | 0.2 | 0.991 | 0.4 | Commissioning of new linacs
27 | 3 | NHS | 2006-05-11 | MeV | 4 | 2003 | -0.07 |  | 1.006 | 0.2 | 1.006 | 0.4 | Measurements at Zref. Zref -> Zmax factor = 1.008
27 | 3 | NHS | 2006-05-11 | MeV | 6 | 2003 | -0.09 |  | 1.004 | 0.2 | 1.004 | 0.4 | Measurements at Zref. Zref -> Zmax factor = 1.001
27 | 3 | NHS | 2006-05-11 | MeV | 8 | 2003 | -0.11 |  | 1.005 | 0.2 | 1.005 | 0.4 | Measurements at Zref. Zref -> Zmax factor = 0.999
27 | 3 | NHS | 2006-05-11 | MeV | 10 | 2003 | -0.08 |  | 1 | 0.2 | 1 | 0.4 | Measurements at Zref. Zref -> Zmax factor = 1
27 | 3 | NHS | 2006-05-11 | MeV | 12 | 2003 | -0.03 |  | 0.999 | 0.2 | 0.999 | 0.4 | Measurements at Zref. Zref -> Zmax factor = 1.001
27 | 3 | NHS | 2006-05-11 | MeV | 15 | 2003 | -0.02 |  | 1.002 | 0.2 | 1.003 | 0.4 | Measurements at Zref, Mean of NACP and ROOS. Zref -> Zmax factor = 1.001
27 | 3 | NHS | 2006-05-11 | MeV | 18 | 2003 | 0.14 |  | 1.001 | 0.2 | 1.001 | 0.4 | Measurements at Zref. Zref -> Zmax factor = 0.992
27 | 3 | NHS | 2009-09-10 | MV | 10 |  | 1.007 | 0.2 | 0.994 | 0.2 | 0.989 | 0.4 | Commissioning of new linacs
8 | 6 | NHS | 2010-02-19 | MV | 6 |  | 0.995 | 0.2 | 0.995 | 0.1 | 1.002 | 0.4
28 | 3 | NHS | 2009-05-11 | MeV | 6 | 2003 |  |  | 0.994 | 0.2 | 0.992 | 0.4
28 | 3 | NHS | 2009-05-11 | MeV | 9 | 2003 |  |  | 0.991 | 0.2 | 0.99 | 0.4
28 | 3 | NHS | 2009-05-11 | MeV | 12 | 2003 |  |  | 0.992 | 0.2 | 0.995 | 0.4
28 | 3 | NHS | 2009-05-11 | MeV | 15 | 2003 |  |  | 0.993 | 0.2 | 0.999 | 0.4
28 | 3 | NHS | 2009-05-11 | MeV | 18 | 2003 |  |  | 0.998 | 0.2 | 1.006 | 0.4
28 | 3 | NHS | 2009-05-11 | MeV | 22 | 2003 |  |  | 0.995 | 0.2 | 1.005 | 0.4
8 | 6 | NHS | 2010-02-19 | MV | 10 |  | 0.998 | 0.2 | 0.995 | 0.1 | 1.002 | 0.4
45 | 5 | NHS | 2010-06-24 | MV | 6 |  | 0.992 | 0.15 | 1.001 | 0.1 | 0.998 | 0.4
45 | 5 | NHS | 2010-06-24 | MV | 10 |  | 0.995 | 0.15 | 1.004 | 0.1 | 0.999 | 0.4
30 | 3 | NHS | 2011-05-13 | MeV | 6 | 2003 | 0 |  | 0.993 | 0.2 | 0.994 | 0.4
30 | 3 | NHS | 2011-05-13 | MeV | 8 | 2003 | -0.1 |  | 0.994 | 0.2 | 0.995 | 0.4
30 | 3 | NHS | 2011-05-13 | MeV | 10 | 2003 | -0.1 |  | 0.995 | 0.2 | 0.994 | 0.4
30 | 3 | NHS | 2011-05-13 | MeV | 12 | 2003 | -0.1 |  | 0.991 | 0.2 | 0.996 | 0.4
30 | 3 | NHS | 2011-05-13 | MeV | 15 | 2003 | -0.1 |  | 0.992 | 0.2 | 0.997 | 0.4
30 | 3 | NHS | 2011-05-13 | MeV | 18 | 2003 | -0.1 |  | 0.992 | 0.2 | 0.995 | 0.4
24 | 3 | NHS | 2010-09-08 | MV | 6 |  | 0.995 | 0.2 | 1 | 0.1 | 0.995 | 0.4
24 | 3 | NHS | 2010-09-08 | MV | 10 |  | 0.995 | 0.2 | 0.999 | 0.1 | 0.998 | 0.4 | Watermark saying "DRAFT"
43 | 5 | NHS | 2007-07-05 | kV | 50 |  | 1.077 | 2 | 1.005 | 0.6 | 0.994 | 0.4
43 | 5 | NHS | 2007-07-05 | kV | 120 |  | 1.018 | 2 | 1 | 0.6 | 0.995 | 0.4
43 | 5 | NHS | 2007-07-05 | kV | 260 |  | 0.972 | 2 | 1 | 0.6 | 1.002 | 0.4
42 | 4 | Private | 2007-07-11 | kV | 80 |  | 1.028 | 2 | 1.006 | 0.6 | 1.005 | 0.4
42 | 4 | Private | 2007-07-11 | kV | 95 |  | 1.049 | 2 | 1.018 | 0.6 | 1.011 | 0.4
42 | 4 | Private | 2007-07-11 | kV | 120 |  | 0.993 | 2 | 1.015 | 0.6 | 1.008 | 0.4
42 | 4 | Private | 2007-07-11 | kV | 140 |  | 1.011 | 2 | 1.013 | 0.6 | 1.009 | 0.4
35 | 8 | NHS | 2007-10-01 | kV | 144 |  | 1.04 | 5 | 1.003 | 0.6 | 0.994 | 0.4
30 | 3 | NHS | 2011-05-13 | MV | 6 |  | 0.992 | 0.15 | 0.997 | 0.2 | 1.001 | 0.4 | Measured output data (tabulated gave output ratio of 1.005)
30 | 3 | NHS | 2011-05-13 | MV | 10 |  | 0.995 | 0.15 | 0.99 | 0.2 | 1 | 0.4 | Measured data. Had to manually edit yhe spreadsheet. Copied some numbers between cells and numbers seem senisible.
34 | 8 | NHS | 2001-09-28 | MeV | 6 | 1996 |  |  | 1.012 | 0.5
34 | 8 | NHS | 2001-09-28 | MeV | 12 | 1996 |  |  | 1.01 | 0.5
35 | 8 | NHS | 2013-07-23 | MeV | 6 | 2003 | 0 |  | 1.001 | 0.2 | 1.002 | 0.4 | Calculations of 1996 and 2003 CoPs done and the ratio between each is given in the report. This is approx 3%. IT was founf that the host was using a thermetical Fion value, rather than using 2 voltage method.
35 | 8 | NHS | 2013-07-23 | MeV | 8 | 2003 | 0 |  | 1.004 | 0.2 | 1.005 | 0.4 | Calculations of 1996 and 2003 CoPs done and the ratio between each is given in the report. This is approx 3%. IT was founf that the host was using a thermetical Fion value, rather than using 2 voltage method.
35 | 8 | NHS | 2013-07-23 | MeV | 10 | 2003 | 0 |  | 1.006 | 0.2 | 1.007 | 0.4 | Calculations of 1996 and 2003 CoPs done and the ratio between each is given in the report. This is approx 3%. IT was founf that the host was using a thermetical Fion value, rather than using 2 voltage method.
35 | 8 | NHS | 2013-07-23 | MeV | 12 | 2003 | 0 |  | 1.007 | 0.2 | 1.007 | 0.4 | Calculations of 1996 and 2003 CoPs done and the ratio between each is given in the report. This is approx 3%. IT was founf that the host was using a thermetical Fion value, rather than using 2 voltage method.
35 | 8 | NHS | 2013-07-23 | MeV | 15 | 2003 | 0.1 |  | 0.997 | 0.2 | 0.995 | 0.4 | Calculations of 1996 and 2003 CoPs done and the ratio between each is given in the report. This is approx 3%. IT was founf that the host was using a thermetical Fion value, rather than using 2 voltage method.
35 | 8 | NHS | 2013-07-23 | MeV | 20 | 2003 | -0.1 |  | 1.004 | 0.2 | 0.999 | 0.4 | Calculations of 1996 and 2003 CoPs done and the ratio between each is given in the report. This is approx 3%. IT was founf that the host was using a thermetical Fion value, rather than using 2 voltage method.
35 | 8 | NHS | 2007-10-01 | kV | 225 |  | 1 | 5 | 1.004 | 0.6
3 | 1 | NHS | 2008-10-14 | kV | 100 |  | 1.031 | 2 | 0.998 | 0.6 | 0.999 | 0.4
36 | 8 | NHS | 2007-10-02 | MeV | 6 | 2003 |  |  | 1.008 | 0.2 | 1.008 | 0.4
36 | 8 | NHS | 2007-10-02 | MeV | 9 | 2003 |  |  | 0.997 | 0.2 | 0.997 | 0.4
36 | 8 | NHS | 2007-10-02 | MeV | 12 | 2003 |  |  | 1.005 | 0.2 | 1.005 | 0.4
36 | 8 | NHS | 2007-10-02 | MeV | 16 | 2003 |  |  | 1.005 | 0.2 | 1.005 | 0.4
36 | 8 | NHS | 2007-10-02 | MeV | 20 | 2003 |  |  | 0.999 | 0.2 | 0.999 | 0.4
3 | 1 | NHS | 2011-05-23 | MV | 6 |  | 0.998 | 0.2 | 0.995 | 0.2 | 0.995 | 0.4
3 | 1 | NHS | 2011-05-23 | MV | 6 |  | 1 | 0.2 | 0.995 | 0.2 | 0.993 | 0.4 | SRS
3 | 1 | NHS | 2011-05-23 | MV | 10 |  | 0.999 | 0.2 | 0.995 | 0.2 | 0.994 | 0.4
3 | 1 | NHS | 2013-08-28 | MV | 6 |  | 1.002 | 0.2 | 0.993 | 0.2 | 1.001 | 0.4
41 | 4 | Private | 2005-01-06 | MV | 6 |  | 1.003 | 0.15 | 0.998 | 0.1 | 1.004 | 0.4
41 | 4 | Private | 2005-01-06 | MV | 10 |  | 1.003 | 0.15 | 1.003 | 0.1 | 1.009 | 0.4
3 | 1 | NHS | 2011-05-23 | kV | 100 |  | 1.047 | 2 | 1.004 | 0.6 | 1.011 | 0.4
19 | 7 | NHS | 2012-01-13 | kV | 80 |  | 1.008 | 2 | 1.01 | 0.6 | 1.001 | 0.4
19 | 7 | NHS | 2012-01-13 | kV | 140 |  | 1.025 | 2 | 1.02 | 0.6 | 0.993 | 0.4
19 | 7 | NHS | 2012-01-13 | kV | 220 |  | 1.019 | 2 | 1.021 | 0.6 | 1.003 | 0.4
32 | 3 | NHS | 2012-06-21 | kV | 80 |  | 0.948 | 2 | 1.003 | 0.6 | 1.004 | 0.4
32 | 3 | NHS | 2012-06-21 | kV | 150 |  | 0.915 | 2 | 1.01 | 0.6 | 0.996 | 0.4
32 | 3 | NHS | 2012-06-21 | kV | 170 |  | 0.873 | 2 | 0.996 | 0.6 | 1 | 0.4
3 | 1 | NHS | 2013-08-28 | MV | 6 |  | 1.002 | 0.2 | 0.998 | 0.2 | 1.004 | 0.4 | FFF
3 | 1 | NHS | 2013-08-28 | MV | 10 |  | 1.001 | 0.2 | 0.997 | 0.2 | 1.002 | 0.4
3 | 1 | NHS | 2013-08-28 | MV | 10 |  | 1.002 | 0.2 | 1.003 | 0.2 | 1.008 | 0.4 | FFF
3 | 1 | NHS | 2013-08-28 | MV | 15 |  | 1.002 | 0.2 | 0.992 | 0.2 | 0.997 | 0.4
44 | 5 | NHS | 2005-10-11 | MeV | 9 | 2003 |  |  | 0.996 | 0.2 | 1.001 | 0.4
44 | 5 | NHS | 2005-10-11 | MeV | 12 | 2003 |  |  | 1.007 | 0.2 | 1.003 | 0.4
44 | 5 | NHS | 2005-10-11 | MeV | 15 | 2003 |  |  | 1.001 | 0.2 | 1.001 | 0.4
44 | 5 | NHS | 2005-10-11 | MeV | 18 | 2003 |  |  | 0.999 | 0.2 | 1.002 | 0.4
44 | 5 | NHS | 2001-06-04 | MeV | 6 | 1996 |  |  | 1.0029 | 0.5
44 | 5 | NHS | 2001-06-04 | MeV | 9 | 1996 |  |  | 1.0053 | 0.5
3 | 1 | NHS | 2015-02-10 | MV | 6 |  | 1.005 | 0.2 | 0.999 | 0.2 | 1.001 | 0.4
3 | 1 | NHS | 2015-02-10 | MV | 6 |  | 1.005 | 0.2 | 0.999 | 0.2 | 1.002 | 0.4 | FFF
3 | 1 | NHS | 2015-02-10 | MV | 10 |  | 1.002 | 0.2 | 0.999 | 0.2 | 1 | 0.4
3 | 1 | NHS | 2015-02-10 | MV | 10 |  | 1.005 | 0.2 | 0.996 | 0.2 | 0.999 | 0.4 | FFF
45 | 5 | NHS | 2010-06-25 | MeV | 6 | 2003 | 0.02 |  | 1.005 | 0.2 | 1.005 | 0.4 | Data gained from spreadsheets directly
45 | 5 | NHS | 2010-06-25 | MeV | 9 | 2003 | 0.02 |  | 1.008 | 0.2 | 1.001 | 0.4 | Data gained from spreadsheets directly
45 | 5 | NHS | 2010-06-25 | MeV | 12 | 2003 | -0.07 |  | 0.999 | 0.2 | 1.006 | 0.4 | Data gained from spreadsheets directly
45 | 5 | NHS | 2010-06-25 | MeV | 16 | 2003 | -0.05 |  | 1 | 0.2 | 1.007 | 0.4 | Data gained from spreadsheets directly
3 | 1 | NHS | 2015-02-10 | MV | 15 |  | 1.003 | 0.2 | 0.995 | 0.2 | 0.997 | 0.4
47 | 9 | Private | 2011-03-03 | MV | 6 |  | 0.995 | 0.3 | 0.996 | 0.4 | 1.005 | 0.4
47 | 9 | Private | 2011-03-03 | MV | 10 |  | 0.999 | 0.3 | 1.003 | 0.4 | 1.006 | 0.4
47 | 9 | Private | 2011-03-03 | MeV | 4 | 2003 | -0.05 |  | 0.987 | 0.2 | 0.978 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
47 | 9 | Private | 2011-03-03 | MeV | 6 | 2003 | -0.08 |  | 0.985 | 0.2 | 0.983 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
47 | 9 | Private | 2011-03-03 | MeV | 9 | 2003 | -0.12 |  | 0.982 | 0.2 | 0.985 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
47 | 9 | Private | 2011-03-03 | MeV | 12 | 2003 | -0.07 |  | 0.983 | 0.2 | 0.983 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
47 | 9 | Private | 2011-03-03 | MeV | 15 | 2003 | -0.03 |  | 0.983 | 0.2 | 0.986 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
47 | 9 | Private | 2011-03-03 | MeV | 18 | 2003 | -0.12 |  | 0.982 | 0.2 | 0.986 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
48 | 9 | Private | 2013-08-21 | MV | 6 |  | 0.994 | 0.3 | 0.99 | 0.1 | 1.003 | 0.4
48 | 9 | Private | 2013-08-21 | MV | 10 |  | 0.997 | 0.3 | 0.99 | 0.1 | 1.004 | 0.4
48 | 9 | Private | 2013-08-21 | MeV | 4 | 2003 | 0 |  | 1.006 | 0.2 | 1.002 | 0.4
48 | 9 | Private | 2013-08-21 | MeV | 6 | 2003 | 0 |  | 1.007 | 0.2 | 1.009 | 0.4
48 | 9 | Private | 2013-08-21 | MeV | 9 | 2003 | 0.1 |  | 1.004 | 0.2 | 1.006 | 0.4
48 | 9 | Private | 2013-08-21 | MeV | 12 | 2003 | 0 |  | 1.004 | 0.2 | 1.015 | 0.4
48 | 9 | Private | 2013-08-21 | MeV | 15 | 2003 | 0.2 |  | 1.001 | 0.2 | 1.006 | 0.4
49 | 9 | Private | 2013-10-24 | MV | 6 |  | 0.996 | 0.3 | 0.991 | 0.1 | 1 | 0.4 | Measured data recorded. tabulated output ratio was 0.989
49 | 9 | Private | 2013-10-24 | MV | 10 |  | 0.997 | 0.3 | 0.989 | 0.1 | 1.006 | 0.4 | Measured data recorded. tabulated output ratio was 0.991
49 | 9 | Private | 2013-10-24 | MeV | 4 | 2003 | 0.02 |  | 0.996 | 0.2 | 1.012 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
49 | 9 | Private | 2013-10-24 | MeV | 6 | 2003 | 0.04 |  | 0.994 | 0.2 | 1.017 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
49 | 9 | Private | 2013-10-24 | MeV | 9 | 2003 | 0.01 |  | 0.989 | 0.2 | 1.001 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
49 | 9 | Private | 2013-10-24 | MeV | 12 | 2003 | 0.05 |  | 0.989 | 0.2 | 1.002 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
49 | 9 | Private | 2013-10-24 | MeV | 15 | 2003 | 0.07 |  | 0.989 | 0.2 | 1 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
50 | 9 | Private | 2012-06-28 | MV | 6 |  | 0.998 | 0.3 | 0.995 | 0.2 | 1 | 0.4 | Measured data recorded. tabulated output ratio was 1.001
50 | 9 | Private | 2012-06-28 | MV | 10 |  | 0.998 | 0.3 | 0.997 | 0.2 | 1.005 | 0.4 | Measured data recorded. tabulated output ratio was 1.002
50 | 9 | Private | 2012-06-28 | MeV | 4 | 2003 | -0.02 |  | 0.978 | 0.2 | 0.991 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
50 | 9 | Private | 2012-06-28 | MeV | 6 | 2003 | 0.03 |  | 0.989 | 0.2 | 1.008 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
50 | 9 | Private | 2012-06-28 | MeV | 9 | 2003 | -0.09 |  | 0.987 | 0.2 | 1.007 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
50 | 9 | Private | 2012-06-28 | MeV | 12 | 2003 | 0.02 |  | 0.988 | 0.2 | 1.011 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
50 | 9 | Private | 2012-06-28 | MeV | 15 | 2003 | -0.12 |  | 0.988 | 0.2 | 1.008 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
50 | 9 | Private | 2012-06-28 | MeV | 18 | 2003 | -0.07 |  | 0.991 | 0.2 | 1.017 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
51 | 9 | Private | 2014-06-05 | MV | 6 |  | 0.997 | 0.3 | 0.997 | 0.2 | 1.001 | 0.4 | Tabulated output data
51 | 9 | Private | 2014-06-05 | MV | 10 |  | 1 | 0.3 | 0.997 | 0.2 | 0.998 | 0.4 | Tabulated output data
51 | 9 | Private | 2014-06-05 | MeV | 4 | 2003 | 0.03 |  | 0.984 | 0.2 | 0.987 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
51 | 9 | Private | 2014-06-05 | MeV | 6 | 2003 | 0.02 |  | 0.989 | 0.2 | 0.991 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
51 | 9 | Private | 2014-06-05 | MeV | 9 | 2003 | -0.07 |  | 0.988 | 0.2 | 0.996 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
51 | 9 | Private | 2014-06-05 | MeV | 12 | 2003 | -0.04 |  | 1.002 | 0.2 | 1 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
51 | 9 | Private | 2014-06-05 | MeV | 15 | 2003 | -0.11 |  | 0.996 | 0.2 | 0.999 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
52 | 9 | Private | 2009-08-06 | MV | 6 |  | 0.996 | 0.2 | 0.989 | 0.1 | 0.996 | 0.4
52 | 9 | Private | 2009-08-06 | MV | 10 |  | 0.997 | 0.2 | 0.984 | 0.1 | 0.992 | 0.4
52 | 9 | Private | 2009-08-06 | MeV | 4 | 2003 | 0.07 |  | 0.985 | 0.2 | 1.002 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
52 | 9 | Private | 2009-08-06 | MeV | 6 | 2003 | 0.03 |  | 0.987 | 0.2 | 0.998 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
52 | 9 | Private | 2009-08-06 | MeV | 9 | 2003 | 0.08 |  | 0.985 | 0.2 | 1.001 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
52 | 9 | Private | 2009-08-06 | MeV | 12 | 2003 | -0.03 |  | 0.985 | 0.2 | 0.997 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
52 | 9 | Private | 2009-08-06 | MeV | 15 | 2003 | 0 |  | 0.983 | 0.2 | 0.994 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
52 | 9 | Private | 2009-08-06 | MeV | 18 | 2003 | -0.05 |  | 0.982 | 0.2 | 1.001 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
53 | 9 | Private | 2009-12-17 | MV | 6 |  | 1 | 0.3 | 1 | 0.1 | 1.003 | 0.4
53 | 9 | Private | 2009-12-17 | MV | 10 |  | 0.999 | 0.3 | 0.994 | 0.1 | 0.999 | 0.4
53 | 9 | Private | 2009-12-17 | MeV | 4 | 2003 | 0.06 |  | 0.976 | 0.2 | 0.987 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
53 | 9 | Private | 2009-12-17 | MeV | 6 | 2003 | 0.04 |  | 0.98 | 0.2 | 0.993 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
53 | 9 | Private | 2009-12-17 | MeV | 9 | 2003 | 0.18 |  | 0.988 | 0.2 | 1.002 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
53 | 9 | Private | 2009-12-17 | MeV | 12 | 2003 | 0.17 |  | 0.988 | 0.2 | 1.001 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
53 | 9 | Private | 2009-12-17 | MeV | 15 | 2003 | 0.3 |  | 0.988 | 0.2 | 1 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
54 | 9 | Private | 2010-07-01 | MV | 6 |  | 0.996 | 0.3 | 0.991 | 0.1 | 1.004 | 0.4
54 | 9 | Private | 2010-07-01 | MV | 10 |  | 0.997 | 0.3 | 0.997 | 0.1 | 1.008 | 0.4
54 | 9 | Private | 2010-07-01 | MeV | 4 | 2003 | 0.04 |  | 0.984 | 0.2 | 0.999 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
54 | 9 | Private | 2010-07-01 | MeV | 6 | 2003 | -0.02 |  | 0.989 | 0.2 | 1.003 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
54 | 9 | Private | 2010-07-01 | MeV | 9 | 2003 | 0.07 |  | 0.986 | 0.2 | 1 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
54 | 9 | Private | 2010-07-01 | MeV | 12 | 2003 | -0.03 |  | 0.984 | 0.2 | 1.006 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
54 | 9 | Private | 2010-07-01 | MeV | 15 | 2003 | 0.01 |  | 0.983 | 0.2 | 0.999 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
54 | 9 | Private | 2010-07-01 | MeV | 18 | 2003 | 0.03 |  | 0.983 | 0.2 | 1.005 | 0.4 | reports note ongoing investigation into factor to be used with solid water which may make resutls match better.
19 | 7 | NHS | 2003-01-13 | MeV | 4 | 1996 |  |  | 1.01 | 0.5 |  |  | Data from paper report
19 | 7 | NHS | 2003-01-13 | MeV | 6 | 1996 |  |  | 1.009 | 0.5 |  |  | Data from paper report
19 | 7 | NHS | 2003-01-13 | MeV | 10 | 1996 |  |  | 1.004 | 0.5 |  |  | Data from paper report
19 | 7 | NHS | 2003-01-13 | MeV | 18 | 1996 |  |  | 1.007 | 0.5 |  |  | Data from paper report
32 | 3 | NHS | 2012-06-21 | kV | 300 |  | 0.995 | 2 | 0.993 | 0.6 | 1.001 | 0.4